# SSP-4 serial interface protocol

(PRELIMINARY)

| SYNTAX | Variable Definitions | Function | Return / Response |
| --- | --- | --- | --- |
| SSTART | #### none | Establish serial connection | ! |
| SGAINx | where x = 1, 2, or 3 | Select instrument gain where 1 = 100X 2 = 10X 3 = 1X | G=x where x = 100, 10, or 1 |
| SIwxyz | where w,x,y,z are integers 0..9 | Set integration time where the sum is z * 10 msec, y * 100 msec, x * 1 second, z * 10 seconds. | I=x where x = integration time in milliseconds |
| SIt | where t = integration time in milliseconds/10 | Set integration time | I=t where t = integration time in milliseconds |
| SCOUNT | none | Start integration | C=y where y = instrument count output i.e. 0 to 66000 |
| STEMxy | #### where xy = 00 to 40 | Set detector cooler temperature (setpoint in -°C) | T=xy where xy = detector temperature setpoint |
| SFTEMP | #### none | Request detector temperature | F=z where z = detector temperature (-°C) |
| SEXITS | #### none | Terminate serial connection | ##### END |

07.29.2026

/misc/SSP-4 protocol.doc